{"text_format":"2","workbook":"xlsx","n_sheets":9,"sheets":["Introduction","Desktop API","Web API","DICOM","Additional PowerScribe XML Info","GE Notes","Infinitt G3 Notes","McKesson Notes","Phillips iSite Notes"],"not_rendered":{"charts":0,"images":11,"macros":0,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Introduction","index":0,"n_rows":31,"n_cols":2,"rows_kept":31,"cut":false,"hidden":false,"merged":[]}
[]
[null,"Doc. Number:        889649"]
["Product:  PowerScribe® 360 | Reporting","Date:                  9/16/2015"]
[]
[]
[]
["PowerScribe 360 Partner Integrations and Configurations"]
["Client API / WebService API/ DICOM"]
[]
[]
["Nuance strives to provide all available information so clients can make timely decisions relating to client and system level API integrations.\n\nNuance does not certify or approve integrations with specific vendor versions. This is done by the vendor. If you need verification that the PACs or RIS vendor you use, has approved specific PowerScribe 360 Versions with a specific version of their product, please contact your RIS or PACs vendor support team."]
[]
["Configuring Client Integrations in PowerScribe"]
["1. Go to Setup > Sites > PACS section."]
["2. Click the +  icon "]
["3. From the Type drop-down list, select the integration type indicated in the following tables."]
["4. In the Name field, type the name of you PACs / RIS Vendor "]
["5. In the Description field enter a description of your choosing"]
["5. In the Mode field, select Master or Slave as appropriate"]
["6. In the data  field, enter the appropriate entry as shown in the following tables"]
["7. Click the Save icon. Your entry should look similar to the example illustration below."]
{"$rows_empty":8}
["NOTE: Nuance employees are not authorized to adjust or configure the settings on a customer's PACS. The customer, or the customer's PACS/ RIS representative, must make any necessary changes to the PACS configuration or settings.\n"]
["Multiple Sites and Multiple Integrations\n■   If the system has multiple sites and the integration type is master, duplicate the master mode integration under each site. \n\n■   If the integration type is slave do not set up duplicate integrations under each site. Set up all slave integrations under the primary site (first site created). When a user logs in, no matter what site they are working under, the system will read the slave integrations created under the primary site and confirm the workstation is set up with the correct directory structures.\n\n■   If there are multiple integrations on a system, and more than one uses file drop integration such as XML, you will need to ensure that each file drop integration drops files into, or reads files from, separate sub folders under c:\\Nuance on the client workstation. For example, an organization using both GE and Hologic should use the following paths:\n■   C:\\Nuance\\GE\\\n■   C:\\Nuance\\Hologic\\"]

{"sheet":"Desktop API","index":1,"n_rows":41,"n_cols":45,"rows_kept":41,"cut":false,"hidden":false,"merged":[]}
["Partner","AccessNET","AGFA","Allscripts","Aspyra","Amicas","Avreo","Brit","Carestream","Cerner","Clario","CPSI","DR Systems","Epic","Epic ","Epic","Fuji","Fuji","GE","GE","GE ","GE","Hologic","Image Info Systems","Infinitt","Intelerad","Intuitive Imaging","MagView","McKesson","Merge","Medical-Insight","Medicalis","MIM Software","MRS","NovaRad","OnePACS","PenRad","Philips","Primordial","RadNet","RamSoft","ScImage","Sectra","Siemens","TeraRecon"]
["Product Name","MedView","IMPAX","Sunrise Imaging Care","AccessNET MedVIEW","Merge","Interworks",null,null,"RadNet","zVision","ImageLink","Unity","Hyperspace","Radient",null,"Synapse","Synapse","Centricity Extends","Centricity PACs IW","Centricity Powerscribe XML","IDR RA1000 Imagecast","SecurView","iQ-VOICE","Infinitt G3 PACS","InteleViewer","ImageQube\nRational Imaging ",null,"HRS A and D","eFilm","Easy-Viz","Medicalis Professional",null,"MRS","NovaPACS / NovaRIS",null,"PenRad","iSite","Prism","eRAD RIS",null,"PICOM Enterprise","SectraCOM","Syngo"]
["Latest Version","6.9.1","v6.2","v15.2","v6.9.1",null,"v7.Yes",{"$nulls":4},"v16.143","v10+",null,null,"Spring 2008 and Higher","v3.2.1 & v4.0","v4.4","v4.0.Yes",null,null,null,"v6.Yes - v7.Yes","v1.1.0","v3.0.11.4",36990,"14.1.0c",null,"v11.Yes; v12.Yes",null,null,"v.2.4.2.Yes; v6.Yes",null,"v7.0","v807+ / v807+",null,"v6.2.04",null,"v2.Yes - v3.Yes","v42.1",null,"v3.3 build 6"]
["Broadcast IiRadWhere"]
["Broadcast iRadwhere Enabled","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes"]
["Broadcast Port #",9090,9090,9090,9090,9091,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9091,9088,9089,9090,9090,9090,9090,9090,9090,9090,null,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090,9090]
["Workflows Supported"]
["Interrupt",{"$nulls":9},"Yes",null,null,"Yes","Yes","Yes","No","Yes",{"$nulls":8},"Yes",null,"Yes",null,null,"Yes",null,"Yes","No",null,"Yes",null,"Yes","Yes",null,"No"]
["Send to Transcription",{"$nulls":9},"No",null,null,"Yes","Yes","Yes","No","Yes",{"$nulls":10},"Yes",null,null,"Yes",null,"Yes","No",null,"Yes",null,"Yes","Yes",null,"No"]
["Resident",{"$nulls":9},"Yes",null,null,"No","No","No","No","Yes",{"$nulls":10},"Yes",null,null,"Yes",null,"Yes","No",null,"Yes",null,"Yes","Yes",null,"No"]
["Draft",{"$nulls":9},"Yes",null,null,"Yes","Yes","Yes","No","Yes",{"$nulls":10},"Yes",null,null,"Yes",null,"No","No",null,"YEs",null,"Yes","Yes",null,"No"]
["Master Mode"]
["Operates in master Mode","No","Yes","Yes","No","No","No",null,null,null,"Yes","Yes",null,"No","No","No","Yes","Yes","No","Yes","Yes","Yes","Yes",null,"Yes","Yes","No",null,"No","No",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Master Integration Type","No","PowerScribe xml","Uses Broadcast IRadWhere","No","No","No",null,null,null,"Uses Broadcast IRadWhere","Uses Broadcast IRadWhere",null,"No","No","No","RadWhererURL/xml\nor\nPowerScribeXML","Uses Broadcast IRadWhere","No","RadWhereURL/xml","PowerScribeXML",null,"PowerScribeXML",null,"Infinitt ","Intelerad","No",null,"No","No",null,"iRadwhere Port","No","iRadwhere Port","No",null,"iRadwhere Port","RadWhere ActiveYes","Uses iRadwhere Port","No",null,"No","Uses iRadwhere Port","No"]
["Master Value ","No","C:\\Nuance\\Agfa\\Status.xml","N/A","No","No","No",null,null,null,"N/A","N/A",null,"No","No","No","http://PACS1/images?accession={accession] (PSxml) C:\\Nuance\\Fuji\\Status.xml","N/A","No","https://<IP of GEServer>/loadstudy.jsp?user_name={username}&password={password}&rsa=true&accession_number={accession}&nat_access=true","C:\\Nuance\\GE\\Status.xml",null,"C:\\Nuance\\Hologic\\status.xml",null,null,"https://integration.intelerad.com;nuance;nuancefallback;<token_provided_by Intelerad>","No",null,"No","No",null,"N/A","No","N/A","No",null,"N/A","http://localhost:8080","N/A","No",null,"No","N/A","No"]
["Can install PS360 on different PC than PACS","No","No","No","No","No","No",null,null,null,"No","Yes",null,"No","No","No","No","No","No","No","No","Yes","No","No","No","No","No",null,"No","No",null,"Yes","No","No","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["PS360 Launches PACS/RIS Application","No","Yes",null,"No","No","No",null,null,null,"Yes","Yes",null,"No","No","No","No","No","No","Yes",null,"Yes","Yes",null,"No","Yes","No",null,"No","No",null,"No","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["PS360 Closes PACS/RIS Application","No","Yes",null,"No","No","No",null,null,null,"Yes","YEs",null,"No","No","No","No","No","No",null,null,"Yes",null,null,"No","Yes","No",null,"No","No",null,"No","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Auto-Open PACS images when open in 360","No","Yes",null,"No","No","No",null,null,null,"Yes","Yes",null,"No","No","No","Yes","Yes","No","Yes","Yes",null,"Yes",null,"No","Yes","No",null,"No","No",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Auto-Close PACS images when report closed in 360","No","Yes",null,"No","No","No",null,null,null,"Yes","Yes",null,"No","No","No","Yes","Yes","No",{"$nulls":5},"No","Yes","No",null,"No","No",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Mark Exam Dictated / Read","No","Yes","Yes","No","No","No",null,null,null,"Yes","Yes",null,"No","No","No","Yes","Yes","No",null,null,"Yes","No",null,"Yes","Yes","No",null,"No","No",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Multiple Accesion Numbers","No","Yes","Yes","No","No","No",null,null,null,"YEs","Yes",null,"No","No","No","Yes","Yes","No",null,null,null,"Yes",null,"Yes","Yes","No",null,"No","No",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","No",null,"No","Yes","No"]
["Slave Mode"]
["Operates in Slave Mode","Yes","Yes","Yes","Yes","Yes","Yes",null,null,"Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["Slave Integration Type","PowerScribeXML","PowerScribeXML","Uses Broadcast IRadWhere","PowerScribeXML","Amicas Out: PowerScribeXML\nAmicas IN: Radwhere URL/xml\n(Create 2 slave entries)","PowerScribeXML",null,null,"Uses Broadcast IRadWhere","Uses Broadcast IRadWhere","Uses Broadcast IRadWhere",null,"PowerScribeXML","Uses Broadcast IRadWhere","Uses Broadcast IRadWhere","PowerScribeXML","Uses Broadcast IRadWhere","RadWhereURL/XML","Dynamic Imaging","PowerScribeXML","ImageCast GE","PowerScribeXML",null,"PowerScribeXML","Uses Broadcast Iradwhere ","Dynamic Imaging",null,"Horizon Med Imaging","PowerScribeXML",null,"iRadwhere Port","PowerScribeXML","Uses iRadwhre ","PowerScribeXML",null,"Uses iRadwhere ","Uses iRadwhere","Uses iRadwhere","Uses iRadwhere",null,"PowerScribeXML","Uses iRadwhere","PowerScribe XML"]
["Slave Value","C:\\Nuance\\MedView\\Study.xml","C:\\Nuance\\Agfa\\study.xml;C:\\Nuance\\Agfa\\Status.xml;password=<password>\n\nNote: Password provided by Agfa","N/A","C:\\Nuance\\Aspyra\\Study.xml","(OUT) C:\\Nuance\\Merge\\study.xml;C:\\Nuance\\Epic\\status.xml\n(IN) \nC:\\Nuance\\Merge\\test.xml","C:\\Nuance\\Avero\\Study.xml",null,null,"N/A","N/A","N/A",null,"C:\\Nuance\\Epic\\study.xml;C:\\Nuance\\Epic\\status.xml","N/A","N/A","C:\\Nuance\\Fuji\\order.xml;c:\\Nunace\\Fuji\\status.xml","N/A","C:\\Nuance\\GE\\order.xml",111,"C:\\Nuance\\GE\\study.xml;C:\\Nuance\\GE\\status.xml;UseWindowsLogin=True","SAVEONCANCEL=True","C:\\Nuance\\HologiYes\\study.xml;C:\\Nuance\\Hologic\\status.xml;password=holYes;base64;multisite=all",null,"CL\\Nuance\\G3\\study.xml;C:\\Nuance\\G3\\response.xml;password=1234",null,5050,null,"C:\\Nuance\\HRS;multisite=all;Draft=Dictated","C:\\Nuance\\Merge\\study.xml;C:\\Nuance\\Merge\\status.xml;password=merge",null,"N/A","C:\\Nuance\\MIM\\study.xml","N/A","C:\\novarad\\study.xml",null,"N/A","N/A","N/A","N/A",null,"C:\\Nuance\\ScImage\\study.xml;C:\\Nuance\\MScImage\\status.xml","N/A","C:\\Nuance\\Syngo\\boomerang.xml;multisite=All"]
["Can install PS360 on different PC than PACS","Yes","No","No","Yes","No","No",null,null,"Yes","No","Yes",null,"No","No","Yes","No","No","Yes","Yes","Yes","No","No",null,"No","No","Required",null,"Yes","Yes",null,"Yes","No","No","No","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["PACS/RIS Launches PS360 Application","No","Yes",null,"No","Yes","Yes",null,null,"Yes","Yes","Yes",null,"Yes","Yes","Epic  2012 and Higher","No","Yes - User will have to enter their credentials","Yes","No",null,"Yes","Yes",null,"Yes","Yes","No",null,"Yes","Yes",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["PACS/RIS Closes PS360 Application","No","Yes",null,"No","Yes","Yes",null,null,"Yes","Yes","No",null,"Yes","Yes","Epic  2012 and Higher","No","Yes","Yes","No",null,"Yes","Yes",null,"Yes","Yes","Yes",null,"Heavy client HRS only","Yes",null,"Yes","No","Yes","No",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["Auto-open 360 Report when images open","Yes","Yes",null,"Yes","Yes","No",null,null,"Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,null,"Yes",null,"Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["Auto-Close PACS images when report closed in 360","No","Yes",null,"No","Yes","No",null,null,"Yes","Yes","No",null,"Yes","Yes","Yes","Yes","Yes","Yes","No",null,null,"Yes",null,"Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","No","Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"No","Yes","Yes"]
["Mark Eam Dictated / Read",null,"Yes","Yes","Yes","Yes","Yes",null,null,"Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes","No",null,"Yes","Yes",null,"Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["Multiple Accession Numbers",null,"Yes","Yes","No","Yes","Yes",null,null,null,"YEs",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","No",null,null,"Yes",null,"Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","No",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes"]
["Version Approved"]
[]
["PS360 V 1.x",null,"Yes",null,null,"Yes",null,null,"Yes","2012.01.08 to 2012.01.27",null,null,"Yes",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,null,null,"Yes","Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes"]
["PS360 V 1.1",null,"Yes",null,null,"Yes",null,null,"Yes",null,null,null,"Yes",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,null,null,"Yes","Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes",null,"Yes","Yes","Yes","Yes"]
["PS360 V 1.5.x",null,"Yes",null,"Yes","Yes","Yes",null,"Yes",null,null,null,"Yes",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes"]
["PS360 V 2.0.x",null,"Yes",null,"Yes","Yes","Yes",null,"Yes","2012.01.28 to 2012.01.45\nand\n2015.01.01 to 2015.01.09","Yes","Yes","Yes",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes"]
["PS360 V 2.5.x",null,"Yes","Yes","Yes","Yes","Yes",null,"Yes","2012.01.36 to 2012.01.45\n","Yes","Yes","Yes",null,null,"Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes",null,"Yes","Yes","Yes","Yes","Yes","Yes",null,"Yes","Yes"]
["PS360 V3.0.x",null,"No","No","No","No","No","No","No","2012.01.28 to 2012.01.45\nand\n2015.01.01 to 2015.01.09\nand \n2012.01.36 to 2012.01.45\n\n\n","No","No","No","No","No","No","Yes","Yes","No","No","No","No","No","No","No","No","No","No","Yes","No","No","No",null,"No","No","No","No","No","No","No","No","No","No","No","No"]

{"sheet":"Web API","index":2,"n_rows":14,"n_cols":4,"rows_kept":14,"cut":false,"hidden":false,"merged":[[0,0,0,3]]}
["Web Service API"]
[]
["Partner","Product Name","Latest Version","Sol'n Type"]
["ACR.Org"]
["Bayer-Medrad","Certegra","v3.2","contrast dose mgt"]
["Bayer-Radimetrics","Radimetrics",null,"rad dose mgt"]
["ClickView","ClickView 7i","v7.2","Informatics Tools"]
["Fuji","Synapse","v3.2.1 & v4.x","PACS"]
["GE-DoseWatch","DoseWatch","v1.2+","dose mgt"]
["GE-Viewpoint"]
["MRS","MRS","v7.0","Mammo"]
["Radlogics",null,null,"imaging"]
["Siemens - ultrasound",null,null,"ultrasound"]
["Toshiba - WSAPI","Aplio XG/Aplio MX/Xario XG","v5.21/v2.21/v4.21","ultrasound"]

{"sheet":"DICOM","index":3,"n_rows":9,"n_cols":3,"rows_kept":9,"cut":false,"hidden":false,"merged":[[0,0,0,2]]}
["DICOM Integrations"]
[]
["Partner","Product Name","Latest Version"]
["Philips Ultrasound"]
["Siemens Ultrasound"]
["Toshiba Ultrasound","Aplio 300/400/500","v4.x"]
["Vidistar"]
["Lexmark","Modlink"]
["CLICKVIEW","Informatics Tools","4.1.3.0"]

{"sheet":"Additional PowerScribe XML Info","index":4,"n_rows":21,"n_cols":1,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["Additonal PowerScribeXML Integrations Settings"]
["UseWindowsLogin: If specified as true, a single sign-on occurs automatically. The PowerScribe 360 | Reporting user name\nwill be taken from the current Windows user's user name. Note that the PACS/RIS does not need to send any login information\nin the XML. PowerScribe 360 | Reporting attempts to log in every time a request is sent by the PACS/RIS. The example above\n(in the Slave Mode section) uses this option.\nIn addition, for the UseWindowsLogin feature to work properly, you must log in to the PowerScribe 360 | Reporting Administrator\nPortal and select the Allow null password via automation preference, located in Setup > System > Preferences > Security.\n"]
{"$rows_empty":14}
["MultiAccSeperator=\\    - Identifies the character seperating accession unumbers in the acecssion number field in XML File"]
["Multisite=All   - Tells PS360 to serach all sites in system for accession number(s)"]
["thissite     - Tells PS360 to search the current site only for accession number(s)"]
["selected   - Tells PS360 to search only the site selected in the drop down for accession number(s)"]
["FieldMRN  - Identified the field name that the MRN is located in, in the XML file"]

{"sheet":"GE Notes","index":5,"n_rows":39,"n_cols":1,"rows_kept":39,"cut":false,"hidden":false,"merged":[]}
["GE Centricity PACS-IW Integration"]
{"$rows_empty":24}
["NOTE: The PACS administrator or GE Representative must perform the steps in this section."]
["In order to send over a required suffix (termination) at the end of the accession number, you must load an additional file on each"]
["workstation that uses the GE/PowerScribe Dictate function in the Slave Mode."]
["1. Open a text editor application, such as Notepad, and create the following file."]
["2. In the Port= line, use the port number that you specified above for the Slave mode entry, which in the example was 111."]
["3. Save the file using the name Dictation.ini."]
["4. On each workstation that has the GE viewer loaded, go to C:\\Program Files\\Integrad.3\\MIV."]
["5. Copy the Dictation.ini file into the C:\\Program Files\\Integrad.3\\MIV directory."]
["6. Open the desired study in GE PACS-IW."]
["7. Click the Dictate button."]
[]
["GE RA 1000 "]
["NOTE: The PACs administrator or GE representative must perform the following steps.\n1. Install RA1000 and select IDR Extend for the integration.\n2. When prompted for the location of PowerScribe 360 | Reporting server, enter the value below. This can be changed at a\nlater time, if needed, in the extend.properties file.\nhttp://<servername>/PS360ReportingClient/Client/Nuance.PowerScribe360.application\n3. To test proper function, log into RA1000 with the username and password used to log into PowerScribe 360 | Reporting.\nAfter launching RA1000, the PowerScribe 360 | Reporting Client will also launch and log in automatically. This and other\nsettings can be changed in RA1000 under Utilities/Reporting Integration Preferences.\n4. To open an exam in RA1000 and PowerScribe 360 | Reporting, highlight an exam in the RA1000 worklist and choose the\nappropriate option from the IDR drop-down located just below the worklist. You can also see these options in the Exam\nFunctions menu. By default, the IDR API will provide PowerScribe 360 | Reporting with the MRN and Accession Number\n(ris_exam_id) to launch the appropriate exam (in PowerScribe 360 | Reporting).\nFor customers who use the order_nbr as their accession number, use the following platinum property in\nplatinum_custom.properties:\ncom.ge.med.platinum.gui.work.Exam.useOrderNumberInsteadOfRISExamID=false\nThe integration is now complete and can be handed off to the customer for testing.\n \n"]
["Issue where Always On Top not sticking\nAdd this to Platinum_custom.properties:\ncom.ge.med.platinum.gui.WindowOnTopManager.enableJava6OnTopFix=true\n"]

{"sheet":"Infinitt G3 Notes","index":6,"n_rows":2,"n_cols":1,"rows_kept":2,"cut":false,"hidden":false,"merged":[]}
["Infinitt G3"]
["NOTE: The PACs administrator or PACs vendor Representative must perform the following steps.\n1. Four integration files (EncInf.dll, MFSR.ini, MFSR_XML.exe, and PS360.xml) must be installed in the Common folder.\n2. If you have a customized installation path, modify the External_SR and SampleXML path in the MFSR.ini file.\n3. In the Master Mode, hostname must be changed to the actual IIS server name or IP address. INFINITT supports http and\nhttps for the URL call.\n4. In the Master Mode, the Type field is optional when calling the G3 PACS viewer. If the Type is not specified, the default\ntype is S.\n■\nType=S: If the type = S, an accession number should be one of the parameters calling the PACS. Type S launches the\nPACS directly to the images in the Image Viewer. The user can get back to the PACS Worklist and search for other\nstudies.\n■\nType=V: Similar to type S, but type V does not allow the user to get back to the PACS worklist to search for other studies\nfrom PACS. Specific studies can only be launched/loaded from the third-party application.\n5. The Infinitt integration type is Master Mode ONLY. For the Slave Mode, you must select the type PowerscribeXML. (See\nSite PACS Settings: Slave on page 2 of this document.)\n6. This integration definition requires a path for the response file, even if the INFINITT code never uses it.\n"]

{"sheet":"McKesson Notes","index":7,"n_rows":9,"n_cols":1,"rows_kept":9,"cut":false,"hidden":false,"merged":[]}
["McKesson"]
["This step is to be completed by the PACs Admin\nCopy the StartPS360.bat and the StopPS360.bat files into the C:\\Horizon directory on each PowerScribe 360 | Reporting Client Workstation using the McKesson integration.\n (You can get these two files from your Nuance Field Service Representative or from McKesson support.)\nNOTE: The site administrator should coordinate with the McKesson engineers to set a logout delay of at least 120\nseconds. This allows sufficient time for PS 360 to upload profile data and close before the McKesson HMI closes out."]
["Additonal Settings"]
["For some sites, an additional setting can be added to the Mode entry line for Draft=X, where X will be the text set for ResponseStatus in the response XML file sent to the McKesson integration if the report is closed with a Draft state. For some McKesson PACS sites, this will be Draft=Dictated or Draft=Preliminary. If the Draft=X is not entered in the Mode entry line, the ResponseStatus will be set to Suspended if the report is closed in a Draft state.\nThis customization should be determined by customer personnel and/or the PACS vendor according to the site requirements, to ensure the proper XML response.\n"]
[]
[]
[]
[]
["Communicaitons and statuses with PS360 and McKEsson HRS"]

{"sheet":"Phillips iSite Notes","index":8,"n_rows":63,"n_cols":1,"rows_kept":63,"cut":false,"hidden":false,"merged":[]}
["Install and Configure Philips iSite Plug-in"]
["NOTE: The PACS administrator or PACS Representative must perform the following procedure. \nPerform the following steps to set up PowerScribe 360 | Reporting for the iSiteRadiology integration."]
["1.  Log on to iSite Radiology."]
["2.  Click the P icon in the upper-right corner of the window to open the Preferences dialog box."]
["3.  Install the plug-in as either a System or Machine plug-in, but not both:"]
["NOTE: System plug-ins are installed once for all workstations using the iSite server."]
{"$rows_empty":16}
["NOTE: If you do not want the plug-in available on all of the workstations, you must install the plug-in as a Machine plug-in for each workstation that needs access to the PowerScribe 360 | Reporting integration."]
["4. After making your selection (System or Machine plug-in), click Add."]
["The following dialog box opens."]
{"$rows_empty":8}
["5.  Select (place a check mark in) the Enable in iSite Radiology check box. "]
["Note: iSite Enterprise is not supported."]
["6.  Clear (remove the check mark from) the Disable API check box."]
["7.  The plug-in is located on the PowerScribe 360 | Reporting Administrator Portal Web site in a subfolder named pacs/isite. An example dwould be\nhttp:// PS360Server/pacs/iSite/powerscribe360.html"]
["8.  Click OK to add this plug-in. You are returned to the previous dialog box."]
["9.  Click OK."]
[]
["User Setup"]
["For each iSite user:"]
["1.  Log on to iSite and go to the Preferences Dialog window by clicking on the P icon in the upper right corner of the main window."]
["2.  Under User Preferences, select PowerScribe360 Preferences to open the Powerscribe360 Preferences window."]
["3.  Select the preferences you want to use for this site."]
{"$rows_empty":15}
["IMPORTANT: To automatically open studies in iSite when running in Master mode, there is a PowerScribe 360 ! Reporting user option (Tools > Preferences > Reporting) that you must select in order to launch the study when the report is created or opened in PowerScribe 360 ! Reporting."]
["The following illustration shows the Preferences dialog box in the PowerScribe 360 ! Reporting client software."]
["NOTE: The Philips Advanced Visualization Tool plug-in, used in combination with Philips iSite and PowerScribe\n360 | Reporting, was causing Client Workstation lockups. Customers using the Advanced Visualization plug-in should\nobtain Version 7.4.6.510 or later, available from Philips. If the Advanced Visualization plug-in version is an earlier\nversion but is not being used, the customer should remove that plug-in to eliminate lockups as well.\n"]